# AIHE-EX05 — Local Validation Report Template

Primary chapter: 11 — Evidence, Validation, and Real-World Evaluation

Aligned chapters: 3, 8, 11, and 12

Version: V1.0 | Release date: 2026-08-10

# Purpose and status

Use this record to document the intended use, local cohort, exact system version, discrimination, calibration, threshold-specific performance, subgroup findings, workflow, clinical utility, uncertainty, and the resulting local decision. It is an educational and practice-support resource, not a validated clinical, legal, regulatory, or health technology assessment instrument.

# How to use this record

Define the local decision, intended and excluded uses, comparator, population, setting, and evidence period before analysing performance.

Record the exact model, software, interface, threshold, data, and workflow version tested.

Report calibration and threshold-specific consequences in addition to discrimination.

Examine subgroup performance, access, workflow response, downstream capacity, and clinical utility.

Separate observed results, modelled consequences, assumptions, and unresolved uncertainty.

End with a recommendation, conditions, monitoring thresholds, accountable owner, review date, and revalidation triggers.

Do not enter identifiable patient information or uncontrolled institutional data in public copies.

# Validation record

| Section | Field | Response | Owner | Status |
| --- | --- | --- | --- | --- |
| Protocol | Validation question | Does the exact model and workflow support a controlled pilot for adult medical discharges? | Clinical analytics lead | Complete |
| Data | Cohort definition | Adult medical discharges from January to December 2025; exclusions documented for incomplete follow-up. | Clinical analytics lead | Complete |
| Model | Version tested | Model 1.3.2, interface 2.1, threshold 0.24; silent-mode deployment. | Clinical analytics lead | Complete |
| Performance | Discrimination | AUROC 0.78 (95% interval 0.75–0.81); threshold-specific sensitivity and positive predictive value reported separately. | Clinical analytics lead | Complete |
| Performance | Calibration | Calibration-in-the-large +0.03; slope 0.88; overprediction concentrated in the highest-risk group. | Clinical analytics lead | Complete |
| Subgroups | Subgroup findings | Lower sensitivity among patients with incomplete prior-utilization history; access and notes differences require mitigation. | Clinical analytics lead | Complete |
| Limitations | Main limitation | Single-site retrospective cohort, changing discharge workflow, and limited follow-up for patients receiving care elsewhere. | Clinical analytics lead | Complete |
| Decision | Validation recommendation | Proceed to a time-limited pilot with recalibration, subgroup monitoring, case-management capacity limits, and a prespecified stop rule. | Clinical analytics lead | Conditional approval |

# Decision summary

| Decision | Proceed to a controlled pilot with conditions. |
| --- | --- |
| Mandatory safeguards | Named clinical owner; documented human review; subgroup monitoring; privacy and cybersecurity controls; fallback pathway. |
| Principal uncertainty | Transportability after workflow change and incomplete cross-provider follow-up. |
| Monitoring thresholds | Calibration, alert burden, response rate, subgroup sensitivity, case-management capacity, and readmission outcome. |
| Accountable authority | Executive clinical governance committee. |
| Next review | After three months or earlier after a material change or serious incident. |
| Revalidation triggers | Model, threshold, source data, interface, intended-use, population, workflow, price, or material evidence change. |